Mode

Parameters

VM series Flashlight type

LED Ultraviolet Lamp

Operation Manual

2.2 Technical Parameters

|  |  | VM10 | VM30 | VM50 | VM70+ |
| --- | --- | --- | --- | --- | --- |
| Maximum Intensity of Ultraviolet Ray at 38cm Distance |  | ≥10000 uW/cm2 | ≥13000 uW/cm2 | L:≥8000 uW/cm2 | ≥6000 uW/cm2 |
|  |  |  |  | H:≥20000 uW/cm2 |  |
| Irradiated Area at 38cm Distance (≥1200 uW/cm2) |  | ￠63mm | ￠55mm | ￠55mm | ￠180mm |
| Visible Light Illumination |  | 2.91 lux | 3.40 lux | 3.14 lux | 2.20 lux |
| Black Light Filter |  | ● | ● | ● | - |
| Light Source |  | 365nm UV-A LED |  |  |  |
| Number of Beads |  | 1 | 1 | 1 | 3 |
| ICR18650 |  | ● |  |  | ● |
| ICR26650 |  |  | ● | ● |  |
| Battery Capacity |  | 3.7V ≥2600mAh | 3.7V ≥3800mAh | 3.7V ≥3800mAh | 3.7V ≥9000mAh |
| Number of Battery |  | 1 | 1 | 2 | - parallel - battery pack |
| Charging Time |  | 5H | 6H | 7H | 6.5H |
| Working Hours (Continuous) |  | 4H | 5H | L:18H | 4H |
|  |  |  |  | H:8H |  |
| Outline Dimension | Lamp Head Diameter | 28 mm | 40 mm | 62 mm | 72 mm |
|  | Diameter of Light Inlet Window | 22mm | 34mm | 55 mm | 62 mm |
|  | Length | 128 mm | 144 mm | 240 mm | 163 mm |
| Weight (including Battery) |  | 140g | 240g | 520g | 558g |

Note:Product parameters are for reference only

1· Instrument Introduction

VM10, VM30, VM50 and VM70 portable cold light source LED ultraviolet torches are developed by our company for fluorescent magnetic particle and fluorescent penetration inspection and other industrial areas.   Feature all aluminum alloy construction, super high strength and very low power consumption. The products conform the National Pressure-bearing Equipment NDT Standard (NB/T 47013-2015).

2·Range of Application

2.1 Main Application

- Yushi VM Series adopts high-power 365nm Ultraviolet band LED Lamp Beads produced by Japanese NICHIA company, 100% free from UV-B and UV-C element, spectrum is pure.

- The lifetime of Lamp Beads reaches up to 20000 hours, efficient, energy conservation and environmental protection.

- High power output, quick start, reaches the maximum illuminance as soon as switched on.

- Ultra-low power consumption, high efficiency.

- Cold light source makes cold operation, outputs the illuminance lasted and steadily.

- Optical filter is black vitaglass.

- High strength aviation aluminum alloy shell with anodize surface.

- Enhance thermal design, can be used continuous for long time.

- Configure dedicated high speed charger for each lamp.

- Anti-magnetic field interference function makes it possible to work in the strong magnetic field.

2.3 Main Application

- Fluorescence Penetrant Testing

- Fluorescence Magnetic Particle Flaw Detection

- Fluorescence Leak Detection

2.4 Other Application

- Oil Stain, Blot Testing

- Special Gas, Oil and Gas Pipeline Testing

- Fluorescence Reflection Detection for Special Materials

- Criminal Investigation and Legal Medical Expert Survey

- Paper Money Authentication, Jewel Appreciation

- Agriculture, Foodstuff, Medical Treatment and other relevant industries.

- 3. Matters need Attention

- Do not use the charger which is not matched with this   product.

- Keep away from child.

- Do not irradiate eyes and skin.

- Keep away from fire.

- Please store in low temperature, dry and ventilation.

- If long time no use, please take out the battery. It is suggested to charge and discharge every 3 months.

- Optimum charging environment: 18-30℃, humidity 25-62%.

Statement: To dismantle without authorization or man-made damage are outside the scope of warranty.

4. Related Data

The National Pressure-bearing Equipment NDT Standard (NB/T 47013-2015) requires Black Light Lamp / Ultraviolet Lamp as follows:

- Exposure intensity on the workpiece surface must be greater or equal to 1000 uW/cm².

- The wavelength should be 315nm~400nm, and the peak value   wavelength should be 365nm.

- Light source should meet the specification of GB/T5097.

Packing List

- Main Unit

- Charger

- Rechargeable Battery

- Suit Case

- Manual

- Certificate

- UV Protective Glasses (Except VM10)

Company Information

Yushi Instruments

Address: No. 9 Shenbei Road, Shenbei New District, Shenyang City,

Liaoning Province, China

Tel: +86 024-88321301

Website: www.yushindt.com

Worldwide Guaranty Card:

Model: ___________________ S/N Code: ______________

Date of Purchase: __________  Customer Call: _________

Customer Company:___________________________

Note: Products in the warranty period due to normal use of

the fault, can have free maintain with the warranty card.

1                                                     2

3                                                    4